Instructions for Analytical Plus Electronic Balance

Description

The Analytical Plus Electronic Balance is a high precision weighing instrument.  The range of the balance is from 0.01mg to 210 g (full capacity).

Instructions

Leveling the Balance

- The balance has a level indicator in the top left corner of the weighing chamber and two adjustable leveling feet at the rear.  Adjust the leveling feet until the bubble appears in the center of level indicator circle.

Switch Functions

How to Calibrate the Balance (Required Once a Week)

- Make sure the sliding doors to the balance are shut.

- Turn on balance by pressing              and wait until the display shows    .          g

- Hold             until MENU shows on the display

- Press              once and the display will show CAL

- Press              once and the display will show AUTO

- Press            and the display will show -C-

- Wait while balance is calibrated. This will take about 30 seconds.

Do not disturb until display reads 0.00000 g

Operational Instructions

- Press                to rezero the display.

- Open a sliding door and place the sample to be weighed on the pan. Make sure to close the sliding door after placing the sample.

- Wait for the stability indicator (a small s in the bottom left corner of the display window) to appear before reading the weight.

NOTE: If the sample must be held in a container to be weighed, follow these steps to tare the container.

- Place the empty container on the pan until the weight is displayed.

- Press              . The display will blank, then indicates zero. The container's weight is stored in memory.

- Add the sample to the container.  As the sample is added, the weight is displayed. NOTE: The displayed weight will be only the weight of the sample in the container, not the container itself.

- After removing the material and container from the balance, the balance will display the container's weight as a negative number.

- Press               to reset the balance to zero.

# Text boxes

Sends weight data, statistical data, and GLP data to computer/printer. In menus, allows returning to previous menu step.

Selects weighing units functions or options.  In menus, changes to next step or value.

Turns on the balance if the balance is off, zeros the balance. In menu system,   this button is used to accept a choice or enter a submenu.

Stability indicator