M-44 Ellipsometer Procedure for Realignment

- Note:  These instructions are designed as a complement to the Ex-situ instructions already online.  These instructions can be found at:  http://www.jlab.org/ARC/WM/121/ellipsometer.html

Start up:

- Turn on the lamp power source, the Control Module, and the computer in the order listed.

- Press the silver button on the lamp power source to make sure the light turns on.

- Press cancel when prompted for a password.

- Open Wvase32

Alignment and Calibration:

- Select Hardware from the Window menu.

- Select Initialize and enter your name when asked to do so.

- Select Angle of Incidence from the Move menu and enter 90o as the angle.

- Use the micrometer to move the angle to the 90o position.

- Adjust the height micrometer if the stage is in the way of the light beam.

- The light beam should be roughly centered on the aperture of the detector.

- Select Align Sample from the Acquire Data menu.

- If the crosshairs are not centered, make adjustments using an allen wrench.  See Figure 1 for how this should be done.  One of the set screws adjusts in the y direction and the other adjusts the x direction.

- Once everything is centered, press the escape button.

- Now place a standard silicon dioxide sample on the stage.  This can be seen in figure 2.

Figure 1:  Adjusting the detector 									      Figure 2:  Sample stage

- Change the angle of incidence to 85 o and align the sample.  If the light is not directed into the detector aperture, change the height of the arms using the height micrometer.

- This time, use the knobs on the stage sample to do the x-y alignment.  If the gain is really high and the crosshairs will not stop jumping around, the knobs may have to be turned a lot before any fine adjustments can be made.  Keep adjusting until the jumping slows down.

- Change the angle of incidence to 40o and align the sample.

- Change the angle of incidence to 72o and align the sample.  At this point, large adjustments shouldn’t be needed.  If so, start at 85o again and go through steps 15-18 again.

- Return the angle back to 90o so that straight through calibration may be performed.

- Make sure nothing is blocking the path of the beam.

- Select Calibrate System from the Acquire Data menu and choose the Detector System (Straight Through Configuration) and select ok.

- Change the angle of incidence to 72 and make sure it is aligned.  Then select Calibrate System and choose the Coarse calibration.

- Now you can exchange your sample for the standard silicon dioxide sample.

- Align your sample and select Calibrate System and choose the Normal calibration.

Acquiring Data:

- Select dynamic scan from the acquire data menu and select Acquire Single Scan.

- Select Fit from the Window menu and select Normal Fit.

- This will give you the mean square error (MSE), the thickness and the composition.

- If a value for the index of refraction is needed, select the graph window.  Choose the data menu and select current layer (optical constants).  This will give a graph of the optical constants.

- Select File and then output data points to text file.  This will give you a table of all of the optical constants that correspond to the different wavelengths of visible light.