Instructions for TR-Series Analytical Balance

Description

The TR-Series Analytical Balance is a high precision weighing instrument.  The range of the balance is from 0.1mg to 110 g (full capacity).

Instructions

Leveling the Balance

- The balance has a level indicator in the front of the weighing chamber and two adjustable leveling feet at the front of the balance.  Adjust the leveling feet until the bubble appears in the center of level indicator circle.

Switch Functions

External Calibration (required once per week)

- Turn on the machine using the Display On/Off button.

- Select "Mode." (Center Button; it will automatically display "calibrate")

- Press "Enter." (Center Button)

- Choose a weight (20 grams recommended.)

- Open the side or top door of the balance.

- Using gloves and a pair of tweezers, remove the weight from its case and place it on the center of the weighing pan.

- The balance will recognize the mass and automatically calibrate.

- Wait. Once the machine has calibrated the display will show the exact weight of the weight you used. For example, a 20 gram weight will read 20.000.

- Turn off the machine using the Display On/Off Button.  Shut all sliding doors.

Operational Instructions

NOTE: Accuracy and stabilization time are optimized if the sample is placed as close to the center of the pan as possible.

- Place the sample on the pan of the balance. NOTE: Make sure to close the sliding doors for an accurate reading.

- Press the zero key to access the weighing mode. The word WEIGH will appear in both the upper right-hand corner and the lower left-hand corner.

NOTE: If the sample must be held in a container to be weighed, follow these steps to tare the container.

- Place the empty container on the weighing pan.

- Press zero.

- Add the sample to the container and wait for the stability icon to appear.

- The weight of the sample will appear on the display. NOTE: The displayed weight will be only the weight of the material in the container, not the container itself.

# Text boxes

Level

Indicator

Returns the display to zero. Exits  a setup function.

Provide access to balance functions and setup features.

Stability Icon